SECTION D.1

ALKALINE PERSULFATE DIGESTION FOR NITROGEN

& PHOSPHORUS, TOTAL and DISSOLVED

**CEDR Method Codes: TN L01** (Total Nitrogen) **TDN L01** (Total Dissolved Nitrogen)

**TP L04** (Total Phosphorus) **TDP L01** (Total Dissolved Phosphorus)

Scope and Application

This method describes the digestion procedure for the determination of total dissolved nitrogen (TDN) and total dissolved phosphorus (TDP) in fresh and estuarine surface waters by the alkaline persulfate oxidation technique. The method is suitable for the determination of total nitrogen (TN) and total phosphorus (TP) with necessary precautions to ensure that particulates are fully digested.

Typical analytical ranges are 0.05 to 3.0 mg-N/L for total nitrogen and 0.01 to 0.50 mg-P/L for total phosphorus. Analytical ranges may be extended by digesting and analyzing a diluted sample. A higher calibration range for nitrogen is permitted for filtered samples only (i.e., TDN).

Summary of Method

The persulfate oxidation technique for nitrogen in water is performed under heated alkaline conditions, where all organic and inorganic forms of nitrogen are oxidized to nitrate. As the reaction proceeds, NaOH is consumed and the pH drops to < 2.2, which allows the oxidation of all phosphorus compounds to orthophosphate.

ii) An aliquot of digested sample is analyzed for nitrate and orthophosphate using an automated colorimetric method (Sections D.5 & 6.D.8) to produce total nitrogen and total phosphorus concentrations.

Interferences

Some particulate nitrogen compounds in unfiltered samples may be resistant to alkaline persulfate digestion (i.e., refractory N) and yield low total nitrogen results.

Samples preserved with acid will result in low recoveries of nitrogen unless the pH is neutralized to the same pH as the reagents. (USGS 2003.)

Organic carbon reacts with the persulfate oxidation reagent to form carbon dioxide. Concentrations over 150 mg-C/L may deplete the persulfate before all nitrogen compounds are oxidized and cause a low bias for total nitrogen values.

Apparatus and Materials

Autoclave or pressure cooker capable of maintaining 105-120 °C for 55 minutes. (30 minutes if only analyzing nitrogen.)

Glass Digestion Tubes: 30 mL culture tubes with polypropylene liner-less screw-caps or 40 mL vials with Teflon-lined screw-caps. Acid-rinse digestion tubes with HCl to clean. New tubes should be conditioned prior to their first use by filling with persulfate oxidizing reagent and autoclaving at 100-120 °C for 30 minutes.

Reagents and Standards

Reagent Water: Nitrogen-free reagent water.

Borate Buffer Solution: Add approximately 800 mL of reagent water to a two-liter volumetric flask. Quantitatively transfer boric acid (H3BO3) and low-nitrogen (< 0.001% N) sodium hydroxide (NaOH) to the flask. Allow the solution to cool and dilute to 2.0 liters with reagent water. This solution is stable for two months at room temperature.

Persulfate Oxidizing Reagent: Add 400 mL of reagent water to a one-liter volumetric flask. Quantitatively transfer and dissolve low-nitrogen NaOH and low-nitrogen potassium persulfate (K2S2O8 with < 0.001% N) in the flask. Dilute to one-liter with reagent water and store in a glass reagent bottle. Prepare this solution just before use.

Instrument Wash Water: Prepare auto-analyzer wash water by mixing oxidizing reagent and reagent water in a flask, in the same proportion as added to the digestion tubes, e.g., 2:1 (v/v) ratio. Cover with foil and autoclave the solution for at least 30 min. at 100-120°C; cool and add borate buffer (or 3N NaOH if following SM 4500-P J).

v) Calibration Standards: A laboratory may purchase or prepare stock and working standards. The initial calibration verification standard must be purchased or made from a second source.

1. Potassium nitrate (KNO3) for nitrogen (Section 6.D.5): Use primary standard-grade KNO3 that has been oven-dried and desiccated. The primary stock standard is stable for up to 6 months if stored at 4°C.
2. Potassium dihydrogen phosphate (KH2PO4) for phosphorus (Section 6.D.8): Use anhydrous, primary standard grade KH2PO4 that has been oven-dried and desiccated. The primary stock solution is stable for 6 months if stored at 4°C.
3. Prepare a series of working standards just prior to digesting each analytical batch by diluting suitable volumes with reagent or ASW water.
4. Standards should bracket the expected concentrations of samples. Analytical ranges may be extended by digesting and analyzing a diluted sample. However, for total nitrogen determinations in unfiltered samples, the upper calibration range should be limited to 3 mg N/L to ensure that the oxidizing reagent is not consumed prior to reducing organic nitrogen.
5. When analyzing estuarine samples of known salinity it is permissible dilute the working standards and instrument wash water with artificial sea water to match the salinity of the samples. Salinity matching is unnecessary if using a flow injection analyzer or if background correction is built into the instrument.
6. When analyzing samples of varying salinities, it is recommended that the standard curve be prepared in reagent water and Refractive Index corrections be made to the sample concentrations. Refractive Index correction is unnecessary if using a flow injection analyzer or if background correction is built into the instrument.

vi) Digestion Check Standards: A laboratory must analyze one nitrogen and one phosphorus digestion check standard to demonstrate that all compounds containing N and P were completely digested. Standards may be purchased or prepared in the laboratory. Some laboratories use certified reference materials to check the calibration as well as the completeness of digestion. Listed below are various digestion check standards cited in the reference methods.  
  
1. Nitrogen Digestion Check Standards

* 1. Glutamic Acid, Stock and Working Standards − Stock solutions are stable for ten months when preserved with chloroform and stored at 4°C. Prepare the working digestion check standard the day of digestion, diluting to volume with wash water solution.
  2. Nicotinic Acid p-touluenesulfonate – See Standard Methods 4500-P J (APHA 2011).
  3. Glycine – Stable for 6 months at 4°C. (USGS 2005)

1. Phosphorus Digestion Check Standards   
   1. Glycerophosphate (*B*-glycerophosphoric acid-disodium salt-5-hydrate) stock solutions are stable for 10 months when preserved with chloroform and stored at 4°C. Prepare the working digestion check standard the day of digestion.
   2. Adenosine triphosphate – See Standard Methods 4500-P J (APHA 2011).

Procedure

Preparation of Analytical Batch

1. Prepare a series of standard solutions covering the analytical range by diluting either the stock or standard solutions.
2. Add 10.0 mL of sample (or a smaller aliquot diluted to 10.0 mL) to a clean glass test tube or vial.   
     
   Use a wide-bore pipette for taking aliquots of whole-water samples. A magnetic stirrer may be necessary to obtain a representative subsample from samples with high suspended solids.
3. Pipette 10.0 mL of each calibration standard, method blank, LCS, CCV, digestion check standard, etc., into a test tube or vial.
4. Prepare method blanks by adding 10.0 mL of reagent water to test tubes or vials.
5. Dispense 5.0 mL of the persulfate oxidizing reagent to each tube. Immediately cap tightly to prevent volatilization, and then invert twice to mix.
6. Prepare a flask of instrument wash water (mix reagent water and persulfate oxidizing agent in a 2:1 (v/v) ratio) and cover with aluminum foil.

ii) Digestion

1. Autoclave the analytical batch at 120°C for at least 55 minutes on the liquid automatic cycle. If only determining total nitrogen, autoclaving at 100°C for at least 30 minutes is permitted. No volatilization of ammonia occurs after autoclaving.
2. Remove tubes from autoclave and cool to room temperature. Digested samples can be held at room temperature or refrigerated for up to one month before analysis.
3. On the day of analysis, add 1.0 mL of buffer or 3N NaOH solution to each tube and mix.
4. Add buffer or 3N NaOH to the digested wash water, depending on the reagent used.
5. Remove any suspended particles remaining in digests by decantation or filtration prior to colorimetric analyses.

iii) Analysis – See procedures for Nitrate + Nitrite (D.5) and Orthophosphate (D.8)

iv) Calibration and Data Reduction

Prepare a standard curve by plotting the peak heights (or integrated areas) of the digested calibration standards against their known concentrations. Designate a digested method blank as having a concentration of 0.0 mg N/L and 0.0 mg P/L.  
  
**Do not select a curve-fitting function that forces a zero *y*-intercept.** The calibration curve will have a positive *y*-intercept that approximates the baseline-corrected absorbance of the digestion blank.

Quality Control

Filtered samples with results greater than the highest calibration standard must be diluted to a concentration within the calibration range and re-analyzed. Whole water samples with results greater than 3.0 mg N/L must be diluted and re-digested prior to re-analyzing.

Digestion check standards: Must achieve a 90-110% recovery of the known concentration.

Method blanks: see Chapter 6, Section C.

Matrix spike samples: see Chapter 6, Section C.

Laboratory duplicates: see Chapter 6, Section C.

Laboratory control samples: see Chapter 6, Section C.

vii) Method detection limits (MDL): Method detection limits should be established as specified in Chapter 6, Section C.

References

APHA 2012. Standard Methods for the Examination of Water and Wastewater, 22nd ed., *Method 4500-P J-2011, Persulfate Method for Simultaneous Determination of Total Nitrogen and Total*

*Phosphorus* (2012).

APHA 2012. Standard Methods for the Examination of Water and Wastewater, 22nd ed., *Nitrogen Method 4500-N C-2011, Persulfate Method* (2012).

USGS 2003. [USGS WRIR 03–4174](http://nwql.usgs.gov/WRIR-03-4174.shtml): [Methods of Analysis by the U.S. Geological Survey National Water Quality Laboratory-Evaluation of Alkaline Persulfate Digestion as an Alternative to Kjeldahl Digestion for Determination of Total and Dissolved Nitrogen and Phosphorus in Water. Methods I-2650-03 (total dissolved nitrogen and total dissolved phosphorus) and I-4650-03 (total nitrogen and total phosphorus).](http://nwql.usgs.gov/pubs/WRIR/WRIR-03-4174.pdf)

Valderrama, J.C., 1981. The Simultaneous Analysis of Total Nitrogen and Total Phosphorus in Natural Waters: Marine Chemistry, v. 21, p. 109–122.